COLORIMETRIC DETERMINATION OF NITRATE – DIAZO COUPLING METHOD

# Principle of Method:  

Nitrate/Nitrite as total Nitrate in soil and plant samples is determined by colorimetry using a SmartChem Discrete Wet Chemistry analyzer.  Samples are extracted and the resulting solution is analyzed using the Diazo coupling method described here. In this reaction a strongly pink azo dye is formed in acidic solution from the coupling of Sulfanilamide and N-(1-naphthyl) ethylenediamine (NED) by Nitrite.  The reaction proceeds in three stages - nitrate is reduced by cadmium metal to nitrite, which then forms a diazide with Sulfanilamide.  The diazide undergoes a substitution reaction with the aromatic ring of NED, forming the para-substituted napthylene dye.

The reduction followed by diazotization is more widely used than direct spectrophotometric assays of Nitrate as it has the advantage of sensitivity and freedom from interference many other methods lack.  

The colored product is measured at 550 nm and the response is proportional to the concentration of the absorbing species.  It obeys Beer’s law such that:

# , where ε is the molar absorptivity and l the path length.

|  |  | C18H16N5O2 |
| --- | --- | --- |
|  |  | 456.4 g·mol-1 |
|  | ε | 4.0 x 104 M-1·cm-1 |
|  | λmax | 550 nm |

# Instrumentation Used:

SmartChem Discrete Wet Chemistry Analyzer, Model 200.  Manufactured by Westco Scientific, Limited, 2007.  Brookfield, CT, USA

- References:

  - SmartChem 200 Method 375-100E-1, Westco Scientific, April 2007

  - Nitrate and Exchangeable Ammonium Nitrogen. D,G. Maynard and Y.P. Kalra. Soil sampling and Methods of Analysis, M.R. Carter, Ed, Canadian Society of Soil Science.1993 Lewis Publisher  p.25-38

  - Determination of Nitrate-Nitrite Nitrogen by Automated Colorimetry, Method 353.2 Revision 2.0, in Methods for the Determination of Inorganic Substances in Environmental Samples, EPA/600/R-93/100, August 1993.

  - G. Norwitz, P. N. Keliher, Analyst, 1985, 110, 689-694, doi: 10.1039/AN9851000689

  - G. Norwitz, P. N. Keliher, Analyst, 1986, 111, 1033-1037, doi: 10.1039/AN9861101033

# Standards Used:

4.1.	Calibration Standard: Ammonium Nitrate, anhydrous ACS grade NH4NO3.

4.2.	External Reference Standards: Sodium Nitrite and Potassium Nitrate, ACS Grade.